## Slide 1
Masao KURIKI
Graduate School of Advanced Sciences of Matter,
Hiroshima University
Lecture material is made from slides
by L.G. Sukhikh (Thomsk Tech. Univeristy)
ALLPPT.com _ Free PowerPoint Templates, Diagrams and Charts
Medical Applications

## Slide 2: 30000 Accelerators in the world
Courtesy R. Baily

## Slide 3: Accelerators in medicine
Accelerators for treatment
Accelerators for radiopharmaceuticals production
Indirect use (e.g. sterilization)

## Slide 4: Radiation sterilization
E.g. syringes sterilization
BINP, Novosibirsk
Notes: 6 MeV power matters

## Slide 5: Radiopharmaceutical production
Pharmaceuticals for nuclear diagnostics: gamma and positron emitters: F-18, N-15, Tc-99 etc.
Pharmaceuticals for nuclear treatment: beta emitters P-32, Sr-89, etc.
Notes: General scheme

## Slide 6: PET
Notes: General scheme

## Slide 7: Particle accelerators: Cyclotron
In a cyclotron the heavy particles (protons, ions) are accelerated along a spiral trajectory guided inside two evacuated half-cylindrical electrodes (dees) by a uniform magnetic field produced between the pole pieces of a large magnet (1 T).
Notes: General scheme

## Slide 8: Fluorine-18 production
Notes: General scheme

## Slide 9: Fluorine-18 production
Proton energies 11-18 MeV
Currents 50-300 uA
Notes: General scheme

## Slide 10: Radionuclides for nuclear therapy
Notes: General scheme

## Slide 11: Radiation therapy

## Slide 12: Radiation treatment - radiotherapy
>95% of radiation therapy is used for cancer treatment
Notes: General scheme

## Slide 13: Radiation
Ionizing radiation is DANGEROUS!
1 Gy (J/kg) per whole body – radiation disease
7 Gy per whole body – 95% lethal result in a few weeks
20-25 Gy per whole body – lethal result in 30 min
Notes: General scheme

## Slide 14: Radiobiology
Radiation kills cells mainly via DNA damage
Notes: General scheme

## Slide 15: Tumor control probability
TCP and NTCP vs. dose
Minimal complications of normal tissues and organs are desired
B point – 35-70 Gy
Notes: General scheme

## Slide 16: Typical medical accelerator
For a radiotherapist
Black box
Dose distribution
Notes: General scheme

## Slide 17: Dose in water
Notes: General scheme

## Slide 18: Electron machines

## Slide 19: Particle accelerators: Betatron
Betatron is a cyclic induction accelerator in which the electrons are made to circulate in a toroidal vacuum chamber (doughnut) that is placed into a gap between two magnet poles.
Conceptually, the betatron may be considered an analog of a transformer:
  Primary current is the alternating current exciting the magnet.
  Secondary current is the electron current circulating in the doughnut.
Notes: General scheme

## Slide 20: Particle accelerators: Betatron
Betatrons were used for external RT in 70-th, 80-th
Finland
Great Britain
Notes: General scheme

## Slide 21: LINACS
Medical linacs accelerate electrons to kinetic energies from 4 to 25 MeV using microwave radiofrequency fields:
  103 MHz : L band
  2856 MHz: S band
  104 MHz: X band
In a linac the electrons are accelerated following straight trajectories in special evacuated structures called accelerating waveguides.
Notes: General scheme

## Slide 22: LINACS
During the past 40 years medical linacs have gone through five distinct generations, each one increasingly more sophisticated:
  Low energy x rays (4-6 MV)
  Medium energy x rays (10-15 MV) and electrons
  High energy x rays (18-25 MV) and electrons
  Computer controlled dual energy linac with electrons
  Computer controlled dual energy linac with electrons combined with intensity modulation
Notes: General scheme

## Slide 23: LINACS
Linacs are usually mounted isocentrically and the operational systems are distributed over five major and distinct sections of the machine:
  Gantry
  Gantry stand and support
  Modulator
  Patient support assembly
  Control console
Notes: General scheme

## Slide 24: LINACS: configuration
In the simplest and most practical configuration:
  Electron source and the x-ray target form part of the accelerating waveguide and are aligned directly with the linac isocentre obviating the need for a beam transport system.
  Since the target is embedded into the waveguide, this linac type cannot produce electron beams.
Notes: General scheme

## Slide 25: Configuration of modern linacs
Notes: General scheme

## Slide 26: LINACS: configuration
For both electrons and photons (Varian):
Notes: General scheme

## Slide 27: LINACS: configuration
Alpha
magnet
Notes: General scheme

## Slide 28: Configuration of modern linacs
Typical modern dual energy linac, incorporating imaging system and electronic portal imaging device (EPID),
Elekta,
Stockholm
Notes: General scheme

## Slide 29: Configuration of modern linacs
Typical modern dual energy linac, with on board imaging system and an electronic portal imaging device (EPID),
Varian,
Palo Alto, CA, USA
Notes: General scheme

## Slide 30: Configuration of modern linacs
Notes: General scheme

## Slide 31: Electron beam transport
Three systems for electron beam bending have been developed:
90o bending
270o bending
112.5o (slalom) bending
Notes: General scheme

## Slide 32: Clinical X-ray beam
Notes: General scheme

## Slide 33: Clinical X-ray beam
Typical electron pulses arriving on the x-ray target of a linac.
The target is insulated from ground, acts as a Faraday cup, and allows measurement of the electron charge striking the target.
Typical values:
Pulse height: 50 mA
Pulse duration: 2 s
Repetition rate: 100 Hz
Period: 104 s
Notes: General scheme

## Slide 34: Clinical electron beam
Notes: General scheme

## Slide 35: Multileaf collimator
Notes: General scheme

## Slide 36: Conventional EBRT
Rectangular shape
Notes: General scheme

## Slide 37: 3D-CRT
Notes: General scheme

## Slide 38: IMRT and VMAT
Notes: General scheme

## Slide 39: proton machines

## Slide 40: Dose depth distribution
Notes: General scheme

## Slide 41: Spread of Bragg peak
Tumour
Proton energy
70-250 MeV
Notes: General scheme

## Slide 42: Pencil beam scanning
Notes: General scheme

## Slide 43: Single scattering
Notes: General scheme

## Slide 44: Uniform scanning (wobbling)
Notes: General scheme

## Slide 45: Intensity modulation pencil beam
Fast intensity modulation is needed (easier with CW isochronous cyclotrons)
Notes: General scheme

## Slide 46: Main blocks
Accelerator
Energy selection
Beam transport
Gantry
Courtesy M. Schippers
Notes: General scheme

## Slide 47: Building
Notes: General scheme

## Slide 48: Accelerator
Cyclotron: fixed energy, energy modulation is needed, CW
Synchrotron: variable energy, complexity
Notes: General scheme

## Slide 49: Energy degrader (PSI)
Notes: General scheme

## Slide 50: Gantry
Notes: General scheme

## Slide 51: Main manufactures
GE, USA (RP)
Japan (RP+PT)
USA (PT)
Germany (RP)
Canada (RP)
Canada (RP)
Belgium (RP+PT)
USA (PT)
C235
Notes: General scheme

## Slide 52: Conclusions
Roughly 35% of accelerator are used for medical purposes, mainly radiotherapy.
Electron linacs are the main “working horses” that are developed in the direction of precise dose delivery
Proton therapy is widely developed but still to expensive and complicated, mostly because of beam transport and delivery.
Notes: General scheme

## Slide 53
Thank you
for your Attention!